## Slide 1: EVOM Vs Flexi-TEER

## Slide 2: EVOM and Flexi-teer in real system
EVOM: 14.7mV and 250-300 Ohms ( means current isn’t accurate or its doing something in addition for resistance calculations)
Flexi-teer: Took forever to settle (20+ seconds), but went to 2k Ohms eventually
Plugged chop sticks into EVOM or Flexi-TEER with adapter and used them on the same transwell

## Slide 3: On simulated system
System = bread boarded resistor in parallel with cap. R1 = 10k, R2 = 690, expected cell monolayer cap = 1uF/cm^2, 0.33 cm^2 grow surface = total cap of 0.33 uF expected. 12 Hz frequency
Chart: 10k standard Ohms, read res vs cap
| 0 | 0.1 | 0.47 | 1 | 2.2000000000000002
 | 10000 | 10000 | 9990 | 9661 | 8142
Chart: 690 standard Ohms, read res vs cap
| 0 | 0.1 | 0.47 | 1 | 2.2000000000000002 | 4.7 | 10
 | 690 | 690 | 689 | 689 | 689 | 687

## Slide 4: Current thoughts
So far no good explanation on response with chopstick electrodes in real system. Good to very good response in breadboard equivalent circuit.
Electrodes need tested next. Outside of fluctuating values for transwell system (unlikely and no good biological explanation), it’s the only extra variable.
Test media only in transwell chamber. Should be stable.